## Slide 1
Final Presentation on Project on
SMART WHEELCHAIR FOR PARALYZED
Presented by
KNR18EE024 JASIRA T P
KNR18EE034 NAYANA RAMESHAN
KNR18EE046 SHRAVAN SREEDEEP
KNR18EE048 SOORYA P
under the guidance of
Dr. C SREEKUMAR
DEPARTMENT OF ELECTRICAL & ELECTRONICS ENGINEERING
GOVT. COLLEGE OF ENGINEERING KANNUR- 670563
June 2022

## Slide 2: OVERVIEW
Introduction Objectives Literature Review Schematic Diagram Smart Wheelchair Wheelchair Design Eye Tracking Algorithm Home Automation System Simulation of Wheelchair | Eye Tracking Working Model Working of Wheelchair Estimation of Components Limitations Future Scopes Conclusion Reference

## Slide 3: INTRODUCTION
Paralyzed stroke patients are unable to normally communicate with their environment.
Part of their body that is under their control is their eyeballs.
Basic day to day operations are switching on basic devices like fan, bulb etc.
Smart Wheel Chair is the solution that we are proposing for these problems.

## Slide 4: OBJECTIVES
To control the electric wheel chair by tracking eye movements.
To design and implement an electric wheelchair.
To track eye movement and control basic devices like bulb, fan etc.

## Slide 5: LITERATURE REVIEW
[1] Diksha Raj, Anish Kalra, Sreejith Krishna N, Suhriday Roy, Dr. Shreenath K N, “Wagon - The Smart Wheelchair”, International Journal of Emarging Technologies and Innovative Research , vol.8, 2021
It comprises a wheelchair that works on hand gestures, voice command inputs, app inputs(Keypad), and head movements input for locomotion of the differently abled persons
Project can be upgraded in terms of speed and the lag can be overcome by replacing them with more advanced tools and sensors.

## Slide 6: LITERATURE REVIEW
[2] Tan Kian Hou, Yagasena and Chelladurai, “Arduino based voice controlled wheelchair”, First International Conference on Emerging Electrical Energy, Electronics and Computing Technologies, vol. 1432, 2019
Smart motorized voice controlled wheelchair.
Voice command controls the movements of the wheelchair.
Microcontroller used is Arduino Uno.

## Slide 7: LITERATURE REVIEW
[3] Ahmad F. Klaib, Nawaf . Alsrehin, Wasen Y. Melhem, Haneen Bashtawi, Aws A. Magableh, “Eye Tracking Algorithms, Techniques, Tools, and Applications with an Emphasis on Machine Learning and Internet of Things Technologies”, Expert Systems with Applications, vol. 166, 2020.
Machine learning and the Internet of Things are essential factors in the development of eye tracking applications.
This eye tracking techniques have more accurate detection results compared with traditional event-detection methods.

## Slide 8: LITERATURE REVIEW
[4] M.H.A. Sibai, S.A. Manap, “A Study on Smart Wheelchair Systems”, International Journal of Engineering Technology and Sciences, vol. 4, 2015
Presented different smart technologies for wheelchairs.
The human - machine interface and the navigation methods and devices, were the two key areas of concentration.

## Slide 9: SCHEMATIC DIAGRAM
CAMERA MODULE
RASPBERRY PI 4
NODE MCU
ARDUINO UNO
LEFT MOTOR
RIGHT MOTOR
Fig 1. Block Diagram of Smart Wheelchair
RELAY CIRCUIT
NODE MCU
MOTOR DRIVER

## Slide 10: SMART WHEELCHAIR
The whole project is divided into 3 parts:
Wheel Chair
Eye tracking module using Raspberry Pi 4
Home automation system using Node MCU

## Slide 11: SMART WHEELCHAIR
Eye tracking module using Raspberry Pi 4:
  Camera Module is used to capture eye movements.
  The eye movement is tracked using the python eye tracking model using Raspberry Pi 4.
  The position of the pupil is transferred to Node MCU.

## Slide 12: SMART WHEELCHAIR
Wheel Chair
  Wheel chair is also controlled on the basis of the control signal received from Raspberry Pi 4.
  The wheel chair operates at low speed for safety.

## Slide 13: SMART WHEELCHAIR
Home automation system using Node MCU:
  From the position received from Raspberry Pi 4, appropriate decision is taken by the Node MCU.
  Control by Node MCU include the switches of light and fan.

## Slide 14: WHEEL CHAIR DESIGN
During the movement of a wheelchair, there are some resistances that appear.
  Rolling resistance
  Aerodynamic resistance
  Inertial resistance

## Slide 15: Calculation of Rolling Resistance
When an unformulated wheel is wrapped in unformed ground or, respectively, a fully resilient wheel is wound on a fully developed ground, there is no resistance to the movement.
In reality both the ground and the wheel suffer from deformations.
Tr = N×a ------(1)
To balance this torque, it is necessary to develop a torque opposite to this.
This torque equals the product of the pulling force Fr and the dynamic radius of the wheel.

## Slide 16: Calculation of Rolling Resistance
Vertical force (Fz) which expresses the mass of a body in strength equals to:
Fz = m×g -----------(2)
where m is the body mass and g is the acceleration of gravity.
The rolling resistance Fr is described by the following formula:
Fr = Fz × μr ------------(3)
Where μr is the rolling resistance coefficient which is estimated for the wheelchairs with non- pneumatic wheels on concrete, with the value of μr = 0.015.

## Slide 17: Calculation of Rolling Resistance
For the calculation of the wheelchair rolling resistance is taken into account the masses of :
  Wheelchair, m1 = 7kg
  User, m2 = 100kg
  Battery, m3 = 2×2.5kg = 5kg
  Motor, m4 = 2×2.35kg = 4.7kg
  Auxiliary weight, m5 = 200g
  Total weight, M= m1+m2+m3+m4+m5
  = 7+100+5+4.7+0.2 = 116.9kg ≈ 117kg

## Slide 18: Calculation of Rolling Resistance
Fz = M ×g ----------(4)
Fz = 117 × 9.81
Fz = 1147.77 N
M is the total mass in kilogram
Fr = Fz×μr ------------(5)
Fr = 1147.77 × 0.015
Fr = 17.2 N

## Slide 19: Calculation of Rolling Resistance
When climbing a vehicle due to road inclination, a resulting force from the weight of the vehicle appears as shown in Fig 2.
This force opposes the movement of the vehicle when it is on the hill.
Fig 2. Forces in road inclinations.

## Slide 20: Calculation of Rolling Resistance
The ascent resistance FB is calculated from the following formula:
FB=Fz× sin α--------------(6)
Where Fz is vertical force in newtons and α is the angle of the road inclination.
Tilt of the road is the ratio of the height H to the distance L (Fig 1) and expresses the height to be developed at a distance of 100 km:
i = Η/L = tan α ≈ sin α------------(7)
To find the angle α, in order to calculate the ascent resistance, the right formula is the following:
α=tan-1(i)

## Slide 21: Calculation of Rolling Resistance
Thus, for a pavement slope of 10o the vehicle has an ascent resistance equals to:
FB= 1147.77 × Sin10 = 199.31 N
Ft= Fr + FB = 17.2+ 199.31 = 216.51 N
Where Ft is the total traction force which is the sum of ascent resistance FB and the rolling resistance Fr
The torque needed on the wheel to overcome these resistors is calculated by the formula:
Ft = Mk × itot/rt => Mk = Ft × rt / itot
Where Mk is the torque on the wheel in Nm

## Slide 22: Calculation of Rolling Resistance
The rt the radius of the wheel and itot its gear transmission ratio from the motor to the wheel.
The motors have a diameter of 12inch (308mm) with a working voltage of 24V, 250Watt and a fixed gear transmission ratio of 23: 1.
The maximum torque that can be produced by the motor according to its operating diagram is 20Nm and is enough to overcome the climb and roll resistors.
Ft = Mk × itot / rt
Mk = Ft × rt / itot
= 216.15 × 0.154 / 23
= 1.5 Nm

## Slide 23: Calculation of Rolling Resistance
According to standard values maximum speed of motor is 220 rpm and the maximum speed that the wheelchair can develop is 12Km/h.
This is calculated by the following formula:
N = V × 60/2 × π × r
V = N × 2 × π × r / 60
= 216.51 × 2 × π × 0.154/60
V = 3.50 m/s
V = 12.6 Km/h
Where V is the maximum speed that wheelchair can achieve in Km/h, n is motors speed in rpm and r is wheelchair’s wheel radius in m.

## Slide 24: Calculation of Power
Work results when a force acts upon an object to cause a displacement (or a motion) or, in some instances, to hinder a motion.
Three variables are of importance in this definition
  Force
  Displacement
  The extent to which the force causes or hinders the displacement
Each of these three variables find their way into the equation for work.
W=F×χ
Where W is the work in Joules, F is force in N and χ is the displacement in meters

## Slide 25: Calculation of Power
Considering the above mathematical formulas and knowing the velocity it follows that:
P = Ft×Vt
Where P is the power in watts Ft is the traction force in N and Vt is the velocity in m/s.
Theoretically, the power given to the vehicle, when it moves at the maximum speed (3.50m/s) it can develop:
P = Ft×Vt
P= 17.2N × 3.50 m/s P= 60.2 Watt

## Slide 26: Calculation of Power
The needed energy to drive the vehicle for one hour is calculated by the formula:
E= P×t
E = 0.0602KWh
Where E is the energy in KWh, P is power in watt and t is the time in hours.
According to the formula below, the vehicle needs 2.50Ah to move one hour at maximum speed.

## Slide 27: Calculation of Power
E =×V/1000
=0.0602×1000/24
= 2.50Ah
Where E is the energy needed in Ah, Qx is the capacity of the battery in Ah and V is the voltage of the battery used in the experiment equal to 24V.

## Slide 28: EYE TRACKING ALGORITHM
Fig 3. Detection of Pupil in Eye Tracking Algorithm

## Slide 29
Fig 4. Flow Chart of Eye Tracking Program

## Slide 30: HOME AUTOMATION SYSTEM
Fig 5. Home Automation System Using Node MCU and Bluetooth Module

## Slide 31: SIMULATION OF WHEELCHAIR
Simulation of Smart Wheel Chair
Fig 6. Simulation Model of Smart Wheelchair

## Slide 32: SIMULATION OF WHEELCHAIR
Simulation of Smart Wheel Chair
Fig 7. Simulation Model of Smart Wheelchair

## Slide 33: SIMULATION OF WHEELCHAIR
Simulation of Smart Wheel Chair
Fig 8. Simulation Model of Smart Wheelchair

## Slide 34: SIMULATION OF WHEELCHAIR
Simulation of Smart Wheel Chair
Fig 9. Simulation Model of Smart Wheelchair

## Slide 35: SIMULATION OF WHEELCHAIR
Simulation of Smart Wheel Chair
Fig 10. Simulation Model of Smart Wheelchair

## Slide 36: EYE TRACKING
Eye tracking program is developed and its results are shown below.
Fig 11.Straight
Fig 12. Closed

## Slide 37: EYE TRACKING
Eye tracking program is developed and its results are shown below.
Fig 13. Left
Fig 14. Right

## Slide 38: RASPBERRY PI 4 CONFIGURATION
Hardware required:
  Raspberry Pi 4
  RJ45 LAN cable
  32GB SD card and adapter
Software required:
  Raspberry Pi OS
  VNC Viewer and VNC Server
  PuTTY Software

## Slide 39: RASPBERRY PI INTERFACE
Fig 15. Raspbian OS Desktop Interface

## Slide 40: EYE TRACKING PROGRAM
Tested eye tracking program on Raspberry Pi 4
Fig 16. Eye tracking program on Raspberry Pi 4

## Slide 41: WORKING MODEL
Implementation of smart wheelchair
Fig 17. Smart wheelchair

## Slide 42: WORKING MODEL
Hardware setup for eye tracking and wheelchair control
Fig 19. Wheelchair control setup
Fig 18. Circuit inside controller setup

## Slide 43: WORKING MODEL
Conversion to electric wheelchair.
Fig 20. Electric conversion
of wheelchair
Fig 21. Chain and sprocket arrangement

## Slide 44: WORKING MODEL
Camera arrangement in wheelchair
Fig 22. Camera setup

## Slide 45: WORKING MODEL
Implementation of Home Automation System
Fig 23. Home Automation System
Fig 24. Web server interface for Home Automation System

## Slide 46: WORKING MODEL
Fig 23. Control of Home Automation System with external device

## Slide 47: WORKING OF WHEELCHAIR
Home automation by eye tracking.
Fig 24. First bulb control
Fig 25. Second bulb Control

## Slide 48: WORKING OF WHEELCHAIR
Forward movement of wheelchair
Fig 26. Forward Movement

## Slide 49: WORKING OF WHEELCHAIR
Left movement of wheelchair
Fig 27. Left Movement

## Slide 50: WORKING OF WHEELCHAIR
Right movement of wheelchair
Fig 28. Right Movement

## Slide 51: LIMITATIONS
Detection is difficult in dim light.
Chain and sprocket system is not efficient.
Detection is difficult for users wearing spectacles and eye liner.
Breaking is not provided.

## Slide 52: FUTURE SCOPES
Use of night vision cameras helps in dim light instead of normal camera.
Use of hub motors improves the efficiency of the wheelchair.
Breaking and speed control with the help of position sensors in the wheels.

## Slide 53: CONCLUSION
Designed and implemented the electric smart wheelchair.
Wheelchair is controlled by tracking eye movement.
The basic devices like fan, bulb, etc. can be controlled either by eye tracking or by web server.
Limitations of the wheelchair is identified.

## Slide 54: REFERENCE
[1]Diksha Raj, Anish Kalra, Sreejith Krishna N, Suhriday Roy, Dr. Shreenath K N, "Wagon-The Smart Wheelchair", International Journal of Emerging Technologies and Innovative Research, vol. 8, Issue 8, pp. b343-b349, August 2021
[2]D. Das, P. Saha and N. Tabassum, "Real Time Distance and Mobility Detection of Obstacles Using a Smart Multisensor Framework for Visually Impaired People", 2020 IEEE Region 10 Symposium (TENSYMP), pp. 590-593, doi: 10.1109/TENSYMP50017.2020.9230880, 2020
[3] Y. C. Lee and C. M. Lee, "Real-Time Smart Home Surveillance System of Based on Raspberry Pi", 2020 IEEE Eurasia Conference on IOT, Communication and Engineering (ECICE), pp. 72-74, doi: 10.1109/ECICE50847.2020.9301929, 2020

## Slide 55: REFERENCE
[4]Reona Cerejo, Valentine Correia and Neil Pereira, “Eye Controlled Wheelchair based on Arduino Circuit”, International Journal of Technical Research and Applications, vol. 3, Issue 6, pp. 94-98, 2015
[5]Muhammad Azlan Alim, Samsul Setumin, Anis Diyana Rosli, Adi Izhar Che Ani, “Development of a Voice-controlled Intelligent Wheelchair System using Raspberry Pi”, International Conference on Applied Sciences, Information and Technology, vol. 846, 2021
[6]V. U. Rani, J. Sridevi and P. M. Sai, "Web Controlled Raspberry Pi Robot Surveillance", 2021 International Conference on Sustainable Energy and Future Electric Transportation (SEFET), pp. 1-5, doi: 10.1109/SeFet48154.2021.9375666, 2021

## Slide 56: REFERENCE
[7]Fahad Wallam and Muhammad Asif, “Dynamic Finger Movement Tracking and Voice Commands Based Smart Wheelchair”, International Journal of Computer and Electrical Engineering, vol. 3, pp. 497-502 ISSN: 1793-8163 doi:10.7763/IJCEE.2011.V3.368, 2021
[8]Al-Okby, Mohammed & Neubert, Sebastian & Stoll, Norbert & Thurow, Kerstin. “Low-cost Hybrid Wheelchair Controller for Quadriplegias and Paralysis Patients”, Advances in Science, Technology and Engineering Systems Journal, pp. 687-694, doi:10.25046/aj020388, 2017
[9]Luis A. Rivera, Guilherme N. DeSouza, “An Automatic Wheelchair Controlled using Hand Gestures”, IEEE, University of Missouri, April 2010.

## Slide 57: REFERENCE
[10]F. Klaib, N. O. Alsrehin, W. Y. Melhem, H. O. Bashtawi,and A. A. Magableh, “Eye tracking algorithms, techniques, tools, and applications with an emphasis on machine learning and internet of things technologies”, Expert Systems with Applications, vol. 166, pp. 114037–114174, 2021
[11]P. A. Punde, M. E. Jadhav and R. R. Manza, "A study of eye tracking technology and its applications", 2017 1st International Conference on Intelligent Systems and Information Management (ICISIM), pp. 86-90, doi: 10.1109/ICISIM.2017.8122153, 2017
[12]Tan Kian Hou, Yagasena and Chelladurai, “Arduino based voice controlled wheelchair”, First International Conference on Emerging Electrical Energy, Electronics and Computing Technologies, vol. 1432, doi:10.1088/1742-6596/1432/1/012064, 2019

## Slide 58: REFERENCE
[13]M. H. A. Sibai and S. A. Manap, “A study on smart wheelchair systems”, International Journal of Engineering Technology and Sciences, vol. 4, pp. 25–35, doi:10.15282/ijets.4.2015.1.4.1033, 2015
[14]Deepak Kumar, Reetu Malhotra, S.R. Sharma, “Design and Construction of a Smart Wheelchair”, Procedia Computer Science, vol. 172, pp. 302-307, doi:10.1016/j.procs.2020.05.048, 2020
[15]G. Boustany, A. El Deen Itani, R. Youssef, O. Chami and Z. O. Abu-Faraj, "Design and development of a rehabilitative eye-tracking based home automation system," 2016 3rd Middle East Conference on Biomedical Engineering (MECBME), pp. 30-33, doi: 10.1109/MECBME.2016.7745401, 2016

## Slide 59: REFERENCE
[16]Joseph K George, Subin K B, Arun Jose, Hima T, “Eye Controlled Home-Automation For Disables”, IOSR Journal of Electrical and Electronics Engineering, pp. 06-08, 2017
[17]Deepak Kumar Lodhi, Prakshi Vats, Addala Varun, Prashant Solanki, Ritakshi Gupta1, Manoj Kumar Pandey, Rajat Butola, “Smart Electronic Wheelchair Using Arduino and Bluetooth Module”, International Journal of Computer Science and Mobile Computing, vol. 5, pp. 433-438, 2016
[18]Airi Ishizuka, Ayanori Yorozu and Masaki Takahashi, “Driving Control of a Powered Wheelchair Considering Uncertainty of Gaze Input in an Unknown Environment”, Applied Sciences , doi: 10.3390/app8020267, 2018

## Slide 60: REFERENCE
[19]M. A. Eid, N. Giakoumidis and A. El Saddik, "A Novel Eye-Gaze-Controlled Wheelchair System for Navigating Unknown Environments: Case Study with a Person with ALS", in IEEE Access, vol. 4, pp. 558-573, doi: 10.1109/ACCESS.2016.2520093, 2016
[20]Mohammed Hayyan Al Sibai and Sulastri Abdul Manap, “A Study on Smart Wheelchair Systems”, International Journal of Engineering Technology and Sciences, doi:10.15282/ijets.4.2015.1.4.1033, 2015

## Slide 61: THANK YOU

## Slide 62: ESTIMATION OF COMPONENTS
SL No. | Component Name | Quantity | Price (Rs)
1 | 250W, 24V, 13.7A Motor | 2 | 4600
2 | 4 Channel Relay(Peak 30V DC Supply) | 1 | 240
3 | 12V, 7.2Ah Battery | 4 | 3000
4 | 15V,2.2A Battery Charger | 1 | 250
5 | Node MCU | 1 | 500
6 | Ultrasonic Sensors | 4 | 400
7 | DC Wire Socket | 1 | 10
8 | Raspberry Pi 4 | 1 | 5100
9 | Webcam | 1 | 1500
10 | Wheelchair | 1 | 7000
Table 1. Estimation

## Slide 63: ESTIMATION OF COMPONENTS
SL No. | Component Name | Quantity | Price
11 | BTS7960B Bridge Motor Driver | 2 | 1600
12 | Fan | 1 | 220
13 | Sprocket and Chain | 2 | 550
14 | Camera Mount | 1 | 200
Total |  |  | 25170
Table 1. Estimation (Contd..)

## Slide 64: COMPONENTS
Fig 30. Components

## Slide 65: COMPONENTS
Fig 31. Raspberry Pi 4
Fig 32. Arduino Uno

## Slide 66: COMPONENTS
Fig 33. 250W, 24V, 13.7A Motor
Fig 34. 4 Channel Relay(Peak 30V DC Supply)

## Slide 67: COMPONENTS
Fig 35. 12V, 7.2Ah Battery
Fig 36. Battery Charger and Connector

## Slide 68: COMPONENTS
Fig 37. Node MCU
Fig 38. Webcam

## Slide 69: COMPONENTS
Fig 39. Node MCU
Fig 40. Cooling Fan